ТЕХНИЧЕСКОЕ ЗАДАНИЕ

на закупку оборудования для определения индекса прочности горных пород методом точечного нагружения (PLT)

1. Наименование оборудования: Переносная установка (аппарат) для испытаний на точечную нагрузку (Point Load Tester) — 2 шт.

2. Целевое назначение: Проведение оперативных экспресс-испытаний керна в полевых условиях в рамках программ геотехнического бурения. Оборудование должно обеспечивать сбор данных по прочности массива.

3. Технические требования:

- Соответствие международным стандартам: Оборудование и методика проведения испытаний должны строго соответствовать стандарту ASTM D5731 и рекомендациям ISRM (International Society for Rock Mechanics).

- Типы испытаний: Установка должна позволять проводить как осевые (axial), так и диаметральные (diametral) испытания цилиндрических образцов керна.

- Диапазон диаметров керна: Совместимость с керном диаметров PQ, HQ и NQ.

- Конструкция:

  - Жесткая переносная рама с гидравлическим нагрузочным устройством.

  - Наличие стальных конических наконечников стандартной геометрии (согласно ASTM D5731).

  - Цифровой или высокоточный аналоговый манометр (индикатор) с функцией фиксации пиковой (разрушающей) нагрузки.

  - Программное обеспечение позволяющее вычислять площадь скола образца и значение предела прочности на одноосное сжатие и растяжение по данным эксперимента

  - Максимальные габариты используемого образца ДхШхВ, мм  - 100*100*100

  - Масса - не более 75 кг

  - Привод – Механо-гидравлический

  - Максимальное давление, бар – 200

  - Максимальная осевая нагрузка,кН – не менее 75

- Условия эксплуатации: Оборудование должно быть пригодно для работы в полевых условиях высокогорья (~4000 м. н.у.м.) при низких температурах (рудник Кумтор).

4. Комплектация (на каждую единицу):

- Основной блок установки с гидравлическим приводом.

- Защитный экран (для обеспечения безопасности при разрушении образца).

- Кейс для транспортировки и хранения.

- Руководство по эксплуатации на русском и английском языках.

- Действующий сертификат о заводской калибровке.

5. Требования к поставщику:

- Предоставление гарантии на оборудование не менее 12 месяцев.

- Возможность поставки расходных материалов и запасных частей (при необходимости).

TECHNICAL SPECIFICATION

Procurement of Point Load Testing (PLT) Equipment for Rock Strength Index Determination

1. Equipment Description

Portable Point Load Tester (PLT) for determining the Point Load Strength Index of rock samples.

Quantity: 2 units.

Source: Technical specifications provided by the customer.

2. Intended Application

The equipment shall be used for rapid field testing of drill core samples as part of geotechnical drilling programs. The system shall provide reliable data for assessing rock mass strength characteristics.

3. Technical Requirements

3.1 Compliance with Standards

The equipment and testing procedures shall fully comply with:

ASTM D5731 – Standard Test Method for Determination of the Point Load Strength Index of Rock;

ISRM (International Society for Rock Mechanics) recommendations.

3.2 Test Types

The system shall be capable of performing:

Axial point load tests;

Diametral point load tests;

on cylindrical rock core specimens.

3.3 Core Size Compatibility

The equipment shall be suitable for testing the following standard drill core sizes:

NQ;

HQ;

PQ.

3.4 Design and Performance Requirements

The equipment shall include:

A rigid portable frame with a hydraulic loading system;

Standard steel conical platens/tips complying with ASTM D5731 requirements;

A digital or high-accuracy analog gauge equipped with a peak load hold function for recording the failure load;

Software capable of calculating:

fracture surface area;

uniaxial compressive strength (UCS);

tensile strength,

based on the test results;

Maximum specimen dimensions: 100 × 100 × 100 mm;

Total equipment weight: not more than 75 kg;

Drive type: mechanical-hydraulic;

Maximum hydraulic pressure: 200 bar;

Maximum axial load capacity: minimum 75 kN.

3.5 Operating Conditions

The equipment shall be suitable for operation in high-altitude field conditions (approximately 4,000 m above sea level) and at low ambient temperatures typical of the Kumtor Mine, Kyrgyz Republic.

4. Scope of Supply (Per Unit)

Each unit shall be supplied complete with:

Main testing unit with hydraulic drive;

Safety shield/guard for operator protection during specimen failure;

Transport and storage case;

Operating manual in both Russian and English;

Valid factory calibration certificate.

5. Supplier Requirements

The supplier shall provide:

A minimum 12-month warranty on the equipment;

Availability of spare parts and consumables supply, if required.